Development and Single Laboratory Validation of Ultra Fast Liquid Chromatography Method for Quantification of Spinetoram content in Suspension Concentrate (SC)                   Formulation

Shubham Yadav1, Samsul Alam*2, Ajin S. Anil3 Lalitesh K. Thakur4, Jitendra Kumar5

Institute of Pesticide Formulation Technology, Sector-20, Udyog Vihar, Gurugram-122016, Haryana, India

Corresponding author E mail: yadavshubham3092@gmail.com

Corresponding author postal address: Analytical Division, Institute of Pesticide Formulation Technology, Sector-20, Udyog Vihar, Gurugram-122016, Haryana, India

ABSTRACT

A novel ultra fast liquid chromatography (UFLC) method has been developed and validated for quantification of Spinetoram in suspension concentrate(SC) formulation, using Shimadzu, packed with C18 (250 mm × 4.6 mm, 5 μm) column. Mixture of acetonitrile:methanol (80:20 v/v)-90% and ammonium formate solution(0.1M)-10%  was used as mobile phase. The flow rate was kept 2.0 ml/min and detection was carried out at 250 nm. The linearity of proposed method was investigated in the range of 200-958 PPM (r2=0.999). The percentage recovery found to be 99.95%. The % RSD values for precision study was <1.156 as per modified Horwitz equation as requirements by CIPAC. The developed method was found to be specific, linear, precise and accurate. This method is also useful for quantification of Spinetoram in their single or combination formulated products, environmental samples (soil, water), and agricultural products for pesticide residue analysis.

Keywords: Spinetoram res; Horwitz equation; Method validation, Method development, CIPAC - Collaborative International Pesticides Analytical council, Uncertainty in measurements.

- INTRODUCTION

Spinetoram is an insecticidal mixture of two active neurotoxic constituents of Saccharopolyspora spinosa (Bacci et al., 2016). It is used to control pest insects in stored grain (Vassilakos et al., 2015) and on domestic cats (Paarlberg et al., 2017). Spinetoram is a semi-synthetic bio-insecticide that is a mixture of two components, 3′-O-ethyl-5, 6-dihydro-spinosyn J (major component) and 3′-O-ethyl spinosyn L (minor component), the common names of which are spinetoram J and L, respectively (Park et al., 2012; Rumbos, Dutton, and Athanassiou, 2018; Zhang, Li, & Lamusi, 2019), and it is predominantly employed to control the insect orders Lepidoptera, Diptera, and Thysanoptera (Zhang et al., 2019). As a derivative of a compound of natural origin, spinetoram implies less environmental risk than many systemic and/or synthetic insecticides (DeAmicis et al., 2011). It must be remarked that no study has so far been published to quantify spinetoram in suspension concentrate(SC) or other formulation products, although an ion mobility-based method has been recently published in which the screening detection limits of 280 pesticides in several food matrices, including spinetoram J, L and honey, were determined. (Bauer et al., 2018). Nevertheless, spinetoram is currently registered for several major and specialty crops in many countries around the world, as for example, it has shown promising results as an alternative to traditional grain protectants (Andrić et al., 2019; Rumbos et al., 2018). It has also been evaluated as a substitute for other insecticides, like neonicotinoids, to control resistant thrips (DAmbrosio et al., 2018).   Subsequently, several methods have been proposed to determine spinetoram in other food matrices, mainly grains and vegetables (Fu et al., 2017; Hengel, 2011; Liu et al., 2011; Kamel et al.,  2010; Ko et al., 2016; Kim et al., 2017; Malhat, 2013; Malhat & Abdallah, 2019; Park et al., 2012; Quian et al.,, 2014; Vassilakos, Athanassiou, & Tsiropoulos, 2015a, 2015b; Zhang, Wei, et al., 2019; Zhao et al., 2015). A solvent extraction with dichloromethane (Liu et al., 2011), acetonitrile (Hengel, 2011) or a mixture of acetonitrile and triethylamine (Kamel et al., 2010), followed by a solid-phase extraction (SPE) with C18 and florisil (Kamel et al., 2010), polymeric (Hengel, 2011) or NH2 cartridges (Liu et al., 2011) was often employed. liquid chromatography (LC) in reverse phase mode using columns with a C18 stationary phase was the technique of choice for determining spinetoram in food matrices; meanwhile, mass spectrometry (MS), and especially tandem mass spectrometers (MS/MS), has been predominantly selected for the detection of spinetoram.

Figure1. Structure of Spinetoram

To the best of our knowledge, there is no reported method for quantification of Spinetoram in pesticide formulations. Thus, efforts were made to develop accurate, selective and sensitive method for quantification of Spinetoram in their pesticide formulation using ultra fast liquid chromatographic method. In the current work developed a simple, reliable and reproducible UFLC method, which was duly validated as per the recommended guidelines of CIPAC (Collaborative International Pesticides Analytical Council) by statistical parameters precision, accuracy-recovery, linearity. Uncertainties in measurements were also calculated for the active ingredient. The method has been applied to the estimation of Spinetoram in suspension concentrate (SC) formulation.

II. EXPERIMENTAL

2.1 Materials: Certified Reference material (CRM) of Spinetoram was procured from Sigma Aldrich. The technical grade material of above active ingredient was obtained from market. The analytical standard was prepared by purification of these technical grade materials. The analytical standard was qualified against CRM and calculated purity found as for Spinetoram- 93.11%. This standard was used for further analysis. Sample of Pesticide formulation containing Spinetoram 11.7% suspension concentrate (SC) was procured from market. HPLC grade acetonitrile was purchased from Fischer Scientific, Mumbai (India). Mili-Q (Millipore India Pvt. Ltd) system used to obtain HPLC grade water.

2.2 Instrumentation: The UFLC system used to perform development and validation of this quantification method is of Shimadzu UFLC comprised of a binary solvent pump, Photo Diode array detector (PDA) and auto sampler with lab solutions software.

2.3 Mobile phase preparation: The mobile phase consist of Mixture of A- acetonitrile:methanol (80:20 v/v) and B- ammonium formate solution(0.1M) and 90:10(v/v) ratio. Mobile phase- B was filtered through a 0.45 μm nylon membrane (Millipore Pvt. Ltd, Bengaluru, India) and degassed in an ultrasonic bath.

2.4 Diluent preparation: Mobile phase used as diluent.

2.5 Standard Preparation: The Standard stock solution prepared in 50 ml volumetric flask by dissolving 50 mg of Spinetoram (93.11%) standard in 10 ml of diluent. This solution then sonicated for 10 minutes and diluted to volume with diluent. This standard solution contains 1 mg/ml of Spinetoram.

2.6 Sample Preparation: Sample solution was prepared by taking about 427 mg of Spinetoram 11.7% SC in 50 ml volumetric flask and about 10 ml of respective diluent was added and sonicated for 10 minutes with intermittent shaking. The content was brought back to ambient temperature and diluted to volume with diluent. The sample was filtered through 0.45μm nylon syringe filter.

2.7 Chromatographic condition: Method involves use of Shimadzu- C18 column with length of 250 mm, internal diameter 4.6 mm and 5 μm particle size of stationary phase. The column oven temperature maintained at 25°C throughout the analysis. Different compositions of mobile phase tried in isocratic mode. Mobile Phase-A: Mobile Phase-B (Acetonitrile:Methanol): 0.1M Ammonium formate solution 90:10(v/v) ratio for Spinetoram was selected which gave good resolution. The flow rate was maintained at 2.0 ml/min and detection at 250 nm for Spinetoram was carried out with injection volume of 20μl.

2.8 Initial analysis of sample: Sample was analyzed in accordance with section 2.5-2.7 and results were tabulated in table 1.

Table 1.Results of initial analysis

| Sr. No | Ingredients | Declared active ingredient content | Active ingredient content determined |
| --- | --- | --- | --- |
|  |  | % m/v | % m/v |
| 1. | Spinetoram | 11.7 | 12.00 |

2.9 Calculation:

Active content (% m/v) for Spinetoram

III. RESULTS AND DISCUSSION

3.1 Development and optimization of UFLC Method: In the present work, an analytical method based on UFLC using PDA detector has been developed and validated for the quantification of Spinetoram in suspension concentrate (SC) formulation. The analytical conditions were selected, keeping in mind the different chemical nature of Spinetoram.

The column selection has been done on the basis of back pressure, resolution, peak shape and day to day reproducibility of retention time. After evaluating all these factors, Shimadzu C18 (250 mm x 4.6 mm, 5 μm particle size) column was found to be giving satisfactory results. The selection of mobile phase is based on the chemical structure of spinetoram. Considerably good results were obtained with ammonium formate solution (0.1M) as mobile phase-B. For the selection of organic constituents of mobile phase-A, acetonitrile:methanol was chosen to reduce the longer retention time and to attain good peak shape. Finally the mobile phase composition consisting of in Mobile phase-A (Acetonitrile:Methanol): Mobile phase-B (0.1M Ammonium Formate Solution) in 90:10 (v/v) ratios for Spinetoram. Optimized proportion of mobile phase has shown good resolution for Spinetoram.

4. Method validation

The method validation was carried out as per CIPAC (Collaborative International Pesticides Analytical Council) guideline.

4.1 Specificity: Specificity of the method was determined by injecting mobile phase blank, between the peaks of active ingredients in standard, sample as well as in mobile phase blank. Spinetoram standard and sample solution. Since there was no interference also peak purity was found satisfactory. Refer figure 2-3.

Figure 2.Chromatogram of blank

Figure 3.Chromatogram of standard preparation

4.2 System Suitability: System suitability is integral part of method validation. % RSD of retention times (R.T.) and peak area of five replicate injections of standard solution were less than 1.0 %.( Refer Table 2).

Table 2.System Suitability of standard solution

| Parameters | Results | Limits |
| --- | --- | --- |
| % RSD of retention time | 0.0665 | < 1.0 |
| % RSD of peak area | 0.2182 | < 1.0 |

4.3 Precision: The Precision was evaluated by repeatability. Each level of precision was investigated by five replicate injections of standard solution of Spinetoram with concentration about 1 mg/ml each and 5 different preparations of same sample. Table 3 showing acceptable % RSD values calculated by modified Horwitz equation.

% RSD = <2(1-0.5 log C) × 0.67

Table 3.Acceptable % RSD values calculated by modified Horwitz Equation

| Sr. no. | Compound | % Analyte (m/v) | Analyte Ratio (C) | Acceptable maximum % RSD |
| --- | --- | --- | --- | --- |
| 1. | Spinetoram | 11.7 | 0.117 | 1.85 |

The results of precision study were expressed as % RSD and were tabulated in Table 4.

Table 4.Results of Precision studies

|  | Spinetoram |
| --- | --- |
| Mean (% m/v) | 12.00 |
| % RSD | 0.44 |

4.4 Linearity: The linearity was evaluated by measuring 5 different concentration levels of standard solution of Spinetoram. The linearity curve plotted concentration of standard (PPM) against mean peak areas and the correlation coefficient value was computed. The summary of the parameters shown in Table5.

Figure 4.Linearity of standard solutions

Table 5.Linearity study

| Linearity Range (PPM) | 200- 958 |
| --- | --- |
| Correlation Coefficient (R2) | 0.999 |
| Slope (m) | 1985 |
| Y-intercept (C) | 19476 |

4.5 Accuracy and recovery: Accuracy (% Recovery) of analytical method was determined at four concentration levels by spiking known amount of pure actives in sample. The accuracy was calculated as % of recovery. The mean recovery results were tabulated in Table 6.

Table 6.Results of accuracy study

| Components | Level | Theoretical content of Spinetoram (%) | Spinetoram content determined or obtained (%) | % Mean Recovery | % RSD |
| --- | --- | --- | --- | --- | --- |
| Spinetoram | 17.26% | 15.32 | 15.315 | 99.95 | 1.16 |

5. Uncertainty in measurement (U): Uncertainty of method was measured through the data of uncertainty due to Repeatability, Calibration uncertainty of equipment or glassware, Readability of equipment, CRM purity of concentration, Linearity of calibration curve and Recovery of the analyte. The Combined Relative Uncertainty (Uc) and Expanded Uncertainty (U) were calculated [26]. Refer Table 7

Table 7.Calculated Combined and Expanded Uncertainty

| Components | Mean Value (% m/v) | Combined Relative Uncertainty (Uc) | Expanded Uncertainty (U) (% m/v) |
| --- | --- | --- | --- |
| Spinetoram | 12.00 | 0.035563 | 0.493782 |

IV. CONCLUSION

An accurate, precise, specific and reliable UFLC method has been developed for quantification of Spinetoram in their Suspension Concentrate (SC) formulation. Method validation study showed that the developed method is specific, linear, accurate, and easily reproducible. This method is also useful for quantification of Spinetoram in their single or combination Suspension Concentrate (SC) formulated product with different strengths and different formulation types. This method can also useful for analysis of environmental samples (soil, water), agricultural products for pesticide residue analysis of same actives but required additional extraction procedure and validation. Hence developed method can be adapted to regular quality control analysis of production samples and environmental samples.

V. ACKNOWLEDGMENT

The authors are thankful to Institute of Pesticide Formulation Technology (IPFT), Gurugram, Haryana, India for encouragement and permission for publication.

VI. Funding

The research was conducted without any funding sources.

VII. Conflict of Interest

The authors declare that they have no conflict of interest.

VI. REFERENCES

- Andric G, Kljajic P, Prazic GM, Trdan SBT, Bodroza SM (2019) Effectiveness of spinosad and spinetoram against three Sitophilus species: Influence of wheat endosperm vitreousness. J Stored Prod Res 83: 209–217.

DOI: 10.1016/j.jspr.2019.07.001

- Bacci LD, Lupi SS, Rossaro B (2016) A review of Spinosyns, a derivative of biological acting substances as a class of insecticides with a broad range of action against many insect pests, 40-52.

- Bauer A, Kuballa J, Rohn S, Jantzen E, Luetjohann J (2018) Evaluation and validation of an ion mobility quadrupole time-of-flight mass spectrometry pesticide screening approach. J Sep Sci 41: 2178–2187.  DOI:10.1002/jssc.201701059

- D́Ambrosio DA, Huseth AS, Kennedy GG (2018) Evaluation of alternative mode of action insecticides in managing neonicotinoid-resistant Frankliniella fusca in cotton. Crop Prot 113: 56-63.  DOI:10.1016/j.cropro.2018.07.011

- DeAmicis C, Edwards NA, Giles MB, Harris GH, Hewitson P, Janaway L, Ignatova S (2011) Comparison of preparative reverse phase liquid chromatography and counter current chromatography for the kilogram scale purification of crude spinetoram insecticide. J Chromatogr A 1218: 6122–6127.  DOI:10.1016/j.chroma.2011.06.073

- Fu Q, Huang Y, Fu Q, Liang J, Ma J, Feng X, Li S (2017) Simultaneous determination of spinetoram and its metabolites in rice using liquid chromatography tandem mass spectrometry. Chinese Journal of Pesticide Science 19: 631–637.  DOI:10.16801/j.issn.1008-7303.2017.0085

- Hengel MT (2011) Expanded method development for the determination of pesticides in dried hops by liquid chromatography with tandem mass spectrometry. J Am Soc Brew Chem 69: 121–126.  DOI:10.1094/ASBCJ-2011-0519-01

- Kamel A, Qian Y, Kolbe E, Stafford C (2010) Development and validation of a multiresidue method for the determination of neonicotinoid and macrocyclic lactone pesticide residues in. ilk, fruits, and vegetables by ultra-performance liquid chromatography/MS/MS. J AOAC Int 93: 389–399.

- Kim YS, Yang JY, Jin NY, Yu YM, Lim CH (2017) Residue analysis of spinetoram and spinosad on paprika leaf using the modified QuEChERS pre-treatment methods. Korean J Agric Sci 44: 487–494.  DOI:10.7744/kjoas.20170043

- Ko AY, Kim H, Do JA, Jang J, Lee EH, Ju YJ, Rhee GS (2016) Development of analytical method for determination of spinetoram residues in livestock using LC-MS/MS. Anal Sci Technol 29: 94–103.  DOI:10.5806/AST.2016.29.2.94

- Liu X, Abd EAM, Park JY, Park JH, Cho SK, Shin HC, Shim JH (2011) Determination of spinetoram in leafy vegetables crops using liquid chromatography and confirmation via tandem mass spectrometry. Biomed Chromatogr 25: 1099–1106.  DOI:10.1002/bmc.1577

- Malhat F, Abdallah O (2019) Residue distribution and risk assessment of two macrocyclic lactone insecticides in green onion using micro-liquid-liquid extraction (MLLE) technique coupled with liquid chromatography tandem mass spectrometry. Environ Monit Assess 191: 584.  DOI:10.1007/s10661-019-7752-1

- Paarlberg TJW, Anthony J, Lisa MY, William GR, Daniel ES (2017) Effectiveness and residual speed of flea kill of a novel spot on formulation of spinetoram (Cheristin®) for cats. Parasites & vectors 10(1): 1-8.

- Park KH, Choi JH, Abd AM, Cho SK, Park JH, Kim BM, Shim JH (2012) Determination of spinetoram and its metabolites in amaranth and parsley using QuEChERS-based extraction and liquid chromatography-tandem mass spectrometry. Food Chem 134: 2552–2559.  DOI:10.1016/j.foodchem.2012.04.066

- Park KH, Choi JH, Abd EAM, Cho SK, Park JH, Kim BM, Yang A, Na TW, Rahman MM, Im GJ Shim JH (2012) Determination of spinetoram and its metabolites in amaranth and parsley using QuEChERS-based extraction and liquid chromatography–tandem mass spectrometry. Food Chem 134(4): 2552-2559.

- Quian C, Wu Q, Lu D, Li J (2014) Using ultra performance liquid chromatography-tandem mass spectrometry to analyze 5 pesticide residues in cowpea. Chinese Journal of Pesticide Science 16: 594–599.  DOI:10.3969/j.issn.1008-7303.2014.05.14

- Rumbos CI, Dutton AC, Athanassiou CG (2018) Insecticidal effect of spinetoram and thiamethoxam applied alone or in combination for the control of major storedproduce beetle species. J Stored Prod Res 75: 56–63.  DOI:10.1016/j.jspr.2017.10.004

- Rumbos CI, Dutton AC, Athanassiou CG (2018) Insecticidal effect of spinetoram and thiamethoxam applied alone or in combination for the control of major stored-product beetle species. J Stored Prod  Res 75: 56-63.

- Vassilakos TN, Athanassiou CG, Tsiropoulos NG (2015a) Persistence and efficacy of spinetoram against three major stored grain beetle on wheat. Crop Prot 69: 44–51.  DOI:10.1016/j.cropro.2014.08.010

- Vassilakos TN, Athanassiou CG, Tsiropoulos NG (2015b) Influence of grain type on the efficacy of spinetoram for the control of Rhyzopertha dominica, Sitophilus granarius and Sitophilus oryzae. J Stored Prod Res 64: 1–7.  DOI:10.1016/j.jspr.2015.02.002

- Vassilakos TN, Christos GA, Nikolaos GT (2015) Influence of grain type on the efficacy of spinetoram for the control of Rhyzopertha dominica, Sitophilus granarius and Sitophilus oryzae. J Stored Prod  Res  64: 1-7.

- Zhang K, Li J (2019) A novel semi-synthesis of spinetoram-J based on the selective hydrolysis of 5, 6-dihydro spinosyn A.  Nat Prod Res 33(19): 2801-2808.

- Zhang Y, Guo W, Chen H, Gao J, Tao L, Li Z, Xu W (2019) The cytotoxic effects of spinetoram on human HepG2 cells by inducing DNA damage and mitochondria-associated apoptosis. Food Agric Immunol 30(1): 1020-1032.

- Zhang Y, Wei J, Long J, Zhu F, Liao G, Duan T, Chen C (2019) A sensitive method for quantification of the insecticide spinetoram in 13 different plant- and animal-type foods by ultra-high performance liquid chromatography-tandem mass spectrometry. International Journal of Environmental Analytical Chemistry 99: 187–198.  DOI:10.1080/03067319.2019.1583336

- Zhao L, Chen G, Zhao J, Zhang Y, Zhu Y, Yang T, Wu YL (2015) Degradation kinetics of the insecticide spinetoram in a rice field ecosystem. Chemosphere 119: 1185–1191.  DOI:10.1016/j.chemosphere.2014.10.026

- National Accreditation Board for Testing and Calibration    Laboratories (NABL) “Guidelines for Estimation and Expression of Uncertainty in Measurement". NABL 141 (2016). http://www.nabl-india.org/nabl